## Sheet1
Bid Ref. B26-07 CT Scan Machine Date:19/03/2026
Required Specifications for: | CT Scan Machine
FDA cleared, year / CE mark, year | Yes
Intended use | Perfusion, Cardiac, Pediatric, Trauma, Neuro, General Radiology
Detector
Crystal material | to be specified
Type of detector | To be specified
No. of detector electronic Channels | to be specified
Detector width, mm | 4 cm
No. of reconstructed slices | 128
Nb. of acquired Slices per rotation | To be specified
Elements per row (Multi elements per row) | to be specified
Spatial reolution | to be specified
GANTRY
Rotation times, sec., 360 degrees | ≤0.35 sec
Gantry tilt, degrees | to be specified
Gantry dimensions, H x W x D, cm | To be specified
Gantry weight, kg | To be specified
Gantry aperture, cm | >70
Scan localizer | Laser
X-RAY TUBE
X-ray Tube model | to be specified
Heat storage, HU | to be specified
Dissipation rate, HU/min | to be specified
Tube cooling | to be specified
Tube focal spot, mm | large 1.1 - 1.2 Small: 0.6-0.7
mA range | to be specified
X-RAY GENERATOR
kW output | >70 KW
kVp range | 80-≥140 KV
mA range | 10 - ≥560
SPIRAL SCANNING
Max scan time, sec | >60 sec
Max scan volume, cm | >195
Pitch | >0.35
Reconstruction time per image, sec | >33 images/sec
PATIENT TABLE
Vertical, cm | to be specified
Longitudinal, cm | >=200
Scannable range, cm | >=200
Max load capacity with accuracy, kg | >220Kg
IMAGE RECONSTRUCTION
Computer CPU | to be specified
Scan FOVs, cm | ≥50 cm
Reconstruction matrices | 1024x1024
Reconstruction time, sec | >33 images/sec
Per slice, sec | as low as 0.03
For localization scan | Real Time
DISPLAY
Monitor size | Dual monitor, to precise the size of each monitor
Matrices, pixels | Medical Grade (hightest Pixels Number)
Max no. of slices displayed at once | to be specified
IMAGE STORAGE
Hard disk, TB | Max TB available
No. online images | to be specified
Archival storage | to be specified
RADIATION DOSE
Dose modulation technique | Must be available
Pediatric-specific dose control | Must be available
Prospective ECG gating | Must be available
Retrospective ECG | Must be available
Low-dose cardiac (axial acquisition) | Must be available
Iterative Reconstruction | Must be available
PERFORMANCE
Min. interscan time, sec | to be specified
Temporal Resolution in msec | to be specified
Temporal resolution in Cardiac Imaging msec | to be specified
Connectivity
DICOM 3.0 interface | Must be available
HIS/RIS Interface | Must be available
MPPS | Must be available
Structured Dose Report | Must be available and tobe sent to PACS
Worklist | Must be available
Other features | - 4D imaging - Motion Correction Technology (Cardiac) - Low dose exposure - High resolution images - Low contrast , Noise free images - Fastest and most versatile scanning - Iterative Model Reconstruction - Fluoroscopy Technology - Perfusion imaging (Neuro, Body) - Cardiac package - Virtual non contrast images - Automatic bone removal - Vessel Analysis - Colon scan - Lung noduls pos-processing - Liver segmentation - TAVI - Fusion (integration and registration) - ECG monitor for the cardiac scanning
Application | Radiation Dose Management
Routine Imaging Application
Cardiac Imaging Application
Pediatric imaging Application
Advanced Imaging Application
Note : Workstation and software license to be quoted and listed seperately
The supplier should provide the shielding calculation and the installation layouts
All clinical features/options must be quoted and listed seperately